## Sheet1
00	No error	elif self.status_hex == '00':self.status_human = 'No error'
01	Process claimed	elif self.status_hex == '01':self.status_human = 'Process claimed'
02	Command error	elif self.status_hex == '02':self.status_human = 'Command error'
03	Process error	elif self.status_hex == '03':self.status_human = 'Process error'
04	Parameter error	elif self.status_hex == '04':self.status_human = 'Parameter error'
05	Parameter type error	elif self.status_hex == '05':self.status_human = 'Parameter type error'
06	Parameter value error	elif self.status_hex == '06':self.status_human = 'Parameter value error'
07	Network not active	elif self.status_hex == '07':self.status_human = 'Network not active'
08	Time-out start character	elif self.status_hex == '08':self.status_human = 'Time-out start character'
09	Time-out serial line	elif self.status_hex == '09':self.status_human = 'Time-out serial line'
0A	Hardware memory error	elif self.status_hex == '0A':self.status_human = 'Hardware memory error'
0B	Node number error	elif self.status_hex == '0B':self.status_human = 'Node number error'
0C	General communication error	elif self.status_hex == '0C':self.status_human = 'General communication error'
0D	Read only parameter.	elif self.status_hex == '0D':self.status_human = 'Read only parameter.'
0E	Error PC-communication	elif self.status_hex == '0E':self.status_human = 'Error PC-communication'
0F	No RS232 connection	elif self.status_hex == '0F':self.status_human = 'No RS232 connection'
10	PC out of memory	elif self.status_hex == '10':self.status_human = 'PC out of memory'
11	Write only parameter	elif self.status_hex == '11':self.status_human = 'Write only parameter'
12	System configuration unknown	elif self.status_hex == '12':self.status_human = 'System configuration unknown'
13	No free node address	elif self.status_hex == '13':self.status_human = 'No free node address'
14	Wrong interface type	elif self.status_hex == '14':self.status_human = 'Wrong interface type'
15	Error serial port connection	elif self.status_hex == '15':self.status_human = 'Error serial port connection'
16	Error opening communication	elif self.status_hex == '16':self.status_human = 'Error opening communication'
17	Communication error	elif self.status_hex == '17':self.status_human = 'Communication error'
18	Error interface bus master	elif self.status_hex == '18':self.status_human = 'Error interface bus master'
19	Timeout answer	elif self.status_hex == '19':self.status_human = 'Timeout answer'
1A	No start character	elif self.status_hex == '1A':self.status_human = 'No start character'
1B	Error first digit	elif self.status_hex == '1B':self.status_human = 'Error first digit'
1C	Buffer overflow in host	elif self.status_hex == '1C':self.status_human = 'Buffer overflow in host'
1D	Buffer overflow	elif self.status_hex == '1D':self.status_human = 'Buffer overflow'
1E	No answer found	elif self.status_hex == '1E':self.status_human = 'No answer found'
1F	Error closing communication	elif self.status_hex == '1F':self.status_human = 'Error closing communication'
20	Synchronisation error	elif self.status_hex == '20':self.status_human = 'Synchronisation error'
21	Send error	elif self.status_hex == '21':self.status_human = 'Send error'
22	Protocol error	elif self.status_hex == '22':self.status_human = 'Protocol error'
23	Buffer overflow in module	elif self.status_hex == '23':self.status_human = 'Buffer overflow in module'

## Sheet2
				complete:-->	dict([(0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c'), (317, 'f'), (318, 'f'), (319, 'i'), (320, 'c'), (321, 'l'), (322, 'c'), (323, 'l'), (324, 'f'), (325, 'f'), (326, 'c'), (327, 'c'), (328, 'c'), (329, 'c')])	dict([(0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity 0%'), (318, 'PIO hardware capacity 100%'), (319, 'Hardware platform id'), (320, 'Hardware platform sub id'), (321, 'Temporary baudrate'), (322, 'Setpoint monitor mode'), (323, 'BHT12'), (324, 'Nominal sensor voltage'), (325, 'Sensor voltage compensation factor'), (326, 'PCB serial number'), (327, 'Minimum measure time'), (328, 'Bus1 parity'), (329, 'Bus2 parity')])	dict([(0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)'), (318, 'PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)'), (319, 'Hardware platform identification number'), (320, 'Hardware platform sub identification number'), (321, 'Temporary volatile Fieldbus2 baudrate (side interface)'), (322, 'Monitor mode for setpoint'), (323, 'Special BHT parameter'), (324, 'Nominal sensor voltage (used for sensor temperature voltage compensation)'), (325, 'Sensor voltage compensation factor (used for sensor temperature voltage comp.)'), (326, 'PCB serial number (unique number or string written by PCB manufacturer)'), (327, 'Minimum measure time between sensors of a piston prover in 0.1 seconds'), (328, 'Fieldbus parity'), (329, 'Fieldbus2 parity (side interface)')])
Number	Process	FB	Var	Name	descr	Var	Name	descr
0	0	0	c	Identification string	identnr.+softwareversion[+serialnr.]	Var, (0, 'c')	Name, (0, 'Identification string')	descr, (0, 'identnr.+softwareversion[+serialnr.]')
1	0	1	c	Primary node address	primary node address: network parameter FLOW-BUS	Var, (0, 'c'), (1, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS')
2	0	2	c	Secondary node address	secundary node address: network parameter FLOW-BUS	Var, (0, 'c'), (1, 'c'), (2, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS')
3	0	3	c	Next node address	next node address: network parameter FLOW-BUS	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS')
4	0	4	c	Last node address	last node address: network parameter FLOW-BUS	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS')
5	0	5	c	Arbitrage	FLOW-BUS arbitrage setting and/or automatic optimization	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization')
6	0	10	c	Initreset	init and reset security key commands for network/parameter settings	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings')
7	1	0	i	Measure	measured value (100% = 32000)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)')
8	1	1	i	Setpoint	setpoint: wanted value (100% = 32000)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)')
9	1	2	i	Setpoint slope	setpoint ramp signal 0..100 % in up to slope x 0.1 sec.	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.')
10	1	3	i	Analog input	analog input signal, normally used for ext. setp. (100% = 32000)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)')
11	1	4	c	Control mode	control mode selection for instrument or module	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module')
12	1	5	f	Polynomial constant A	polynomial constant A (offset)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)')
13	1	6	f	Polynomial constant B	polynomial constant B (span]	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]')
14	1	7	f	Polynomial constant C	polynomial constant C	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C')
15	1	8	f	Polynomial constant D	polynomial constant D	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D')
16	1	9	f	Polynomial constant E	polynomial constant E (offset) for setpoint or power value	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value')
17	1	10	f	Polynomial constant F	polynomial constant F (span) for setpoint or power value	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value')
18	1	11	f	Polynomial constant G	polynomial constant G for setpoint or power value	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value')
19	1	12	f	Polynomial constant H	polynomial constant H for setpoint or power value	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value')
20	1	13	f	Capacity	readout value at 100% in capacity (readout) unit	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit')
21	1	14	c	Sensor type	sensor type information for actual reading and sensor/controller indication	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication')
22	1	15	c	Capacity unit index	pointer to capacity (readout) unit (classic unit table]	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]')
23	1	16	c	Fluid number	pointer to capacity (readout) unit (classic unit table], c	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c')
24	1	17	c	Fluid name	name of fluid	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid')
25	1	18	c	Claim node	node address of module with operation rights	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights')
26	1	19	c	Modify	contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)')
27	1	20	c	Alarm info	status information of several alarms/errors in the instrument	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument')
28	0	12	c	Channel amount	amount of channels which can be operated	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated')
29	0	13	c	First channel	first channel that can be operated	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated')
30	0	14	c	Last channel	last channel that can be operated	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated')
31	9	1	c	<hostcontrl>	operation by HOST computer enable flag	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag')
32	10	0	c	Alarm message unit type	alarm message string with unit type information	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information')
33	10	1	c	Alarm message number	alarm message string with unit number information	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information')
34	10	2	c	Relay status	status of relays/potential free contacts ('0' = not activated, '1' = activated)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
35		0	f	Actual counter value	actual value of counter	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
36		1	c	Signal input selection	signal input selection (' '=no value,'+'=pos value,'-'=neg value input)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
37		2	c	Reset input selection	external reset input enable/disable ('E'=enable,' '=disable)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
38		3	f	<limit>	limit/batch for counter in sensor standard units	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
39		4	c	Delay time	delay time string in days,hours,minutes,seconds	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
40		5	c	Duration time	duration time string in days,hours,minutes,seconds	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
41		6	c	Valve output setting	valve output setting ('0'=do nothing, '1'=close valve)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
42		7	c	Relay output setting	relay output setting ('L'=low, 'H'=high, 'P'=pulse (1 sec.))	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
43		8	c	Operation mode T/A	operation mode of T/A module	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
44		9	c	Readout unit	readout unit string	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
45		10	f	Readout factor	readout factor matching readout unit string	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
46		12	c	Reset unit	reset unit command (1=reset T/A unit)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
47		9	f	Valve differentiator down	valve output differentiation time constant downwards	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
48		10	f	Valve differentiator up	valve output differentiation time constant upwards	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
49		11	f	Sensor differentiator down	sensor signal differentiation time constant downwards	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
50		12	f	Sensor differentiator up	sensor signal differentiation time constant upwards	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)')
51	114	12	c	Cycle time	cycle time * 10 msec. main loop signal processing	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing')
52	115	3	c	Analog mode	analog mode selection for analog operation	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation')
53	116	6	i	Reference voltage	reference voltage output signal for analog operation	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation')
54	114	1	l	Valve output	valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)')
55	117	1	f	Dynamic display factor	dynamic display factor for display filter (0=max, 1=min goes with par 57)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)')
56	117	2	f	Static display factor	static display factor for display filter (0=max, 1=min goes with par 56)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)')
57	115	1	c	Calibration mode	calibration mode selection (not active until cntrlmode has been set to value 9)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)')
58	116	7	i	Valve offset	valve offset: amount of DAC steps within 1 potmeter step	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step')
59	115	2	c	Monitor mode	monitor: output signal (measure) selection for bus and analog output	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output')
60	114	2	c	Alarm register1	alarm register containing warning flags	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags')
61	114	3	c	Alarm register2	alarm register containing critical error flags	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags')
62	116	1	l	<CalRegZS1>	calibration register zero scale input 1 ADC	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC')
63	116	2	l	<CalRegFS1>	calibration register full scale input 1 ADC	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC')
64	116	3	l	<CalRegZS2>	calibration register zero scale input 2 ADC	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC')
65	116	4	l	<CalRegFS2>	calibration register full scale input 2 ADC	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC')
66	114	4	l	ADC control register	ADC control register	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register')
67	116	5	c	Bridge potmeter	sensor bridge zero potmeter setting	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting')
68	115	4	c	<AlarmEnble>	broadcast alarm messsage enable flag	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag')
69	115	5	c	Test mode	test mode selection (not active until cntrlmode has been set to value 5)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)')
70	115	6	c	<ADC channel select>	channel selection ADC	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC')
71	114	5	c	Normal step controller response	controller response for normal steps (128=normal, <128=slower, >128=faster)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)')
72	117	3	f	Setpoint exponential smoothing filter	analog input filter constant (0=max, 1=min)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)')
73	117	4	f	Sensor exponential smoothing filter	sensor input filter constant (0=max, 1=min)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)')
74		21	i	Analog output zero scale	analog output correction factor zero scale (meas outp DSCM-A 0=0 other 32767=0)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)')
75		22	i	Analog output full scale	analog output correction factor full scale (meas outp 2000 = 1 * multiplication)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)')
76		23	i	Analog input zero scale	analog input correction factor zero scale (ext setp DSCM-A 0=0 other 32767=0)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)')
77		24	i	Analog input full scale	analog input correction factor full scale (ext setp 2000 = 1 * multiplication)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)')
78	115	7	c	Tuning mode	(auto)tuning mode selection (not active until cntrlmode has been set to value 6)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)')
79	114	6	c	Valve default	valve type (needed for controlling behaviour)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)')
80	0	19	c	Global modify	contains number(s) of changed processes for indirect polling (0xXX / 0xFF)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)')
81	114	7	f	Valve span correction factor	correction factor valve curve ratio high/low area	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area')
82	114	8	c	Valve curve correction	Valve curve correction for controller (max. factor*0.1, flow where factor = 1)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)')
83	114	9	c	<MemShipNor>	array with memberships for normal Fuzzy controller	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller')
84	114	10	c	<MemShipOpn>	array with memberships for 0-open Fuzzy controller	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller')
85	114	11	c	IO status	IO status byte for jumper settings and LED signal modes	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes')
86	114	13	c	<FuzzStNeNo>	array with neg nor output steps for Fuzzy contr.	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.')
87	114	14	c	<FuzzStPoNo>	array with pos nor output steps for Fuzzy contr.	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.')
88	114	15	c	<FuzzStOpen>	array with open at 0 output steps for Fuzzy contr.	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.')
89	113	1	c	Device type	(FLOW-BUS) device type information string	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string')
90	113	2	c	BHTModel number	model number information string	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string')
91	113	3	c	Serial number	serial number information string (to be changed by Bronkhorst only)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)')
92	113	4	c	Customer model	manufacturing configuration information string	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string')
93	118	1	c	BHT1	special BHT parameter (to be changed by Bronkhorst only)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)')
94	118	2	i	BHT2	special BHT parameter	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter')
95	118	3	l	BHT3	special BHT parameter	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter')
96	118	4	i	BHT4	special BHT parameter	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter')
97	118	5	c	BHT5	special BHT parameter	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter')
98	118	6	c	BHT6	special BHT parameter	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter')
99	118	7	c	BHT7	special BHT parameter	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter')
100	118	8	c	BHT8	special BHT parameter	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter')
101	118	9	l	BHT9	special BHT parameter	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter')
102	118	10	c	BHT10	special BHT parameter (to be changed by Bronkhorst only)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)')
103	114	16	c	Broadcast repeating time	broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)')
104	113	5	c	Firmware version	revision number of firmware	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware')
105	115	9	c	Pressure sensor type	type of pressure sensor	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor')
106	116	8	f	Barometer pressure	mbar atmospheric (central) barometer pressure	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure')
107		25	i	Sensor input zero scale	analog sensor signal input corr. factor zero scale (DSCM-A 0=0 other 32767=0)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure')
108		26	i	Sensor input full scale	analog sensor signal input correction factor full scale (2000=1*multiplication)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure')
109		27	i	Reference voltage input zero scale	analog Vref input correction factor zero scale (DSCM-A 0=0 other 32767=0)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure')
110		28	i	Reference voltage input full scale	analog Vref input correction factor full scale (2000=1*multiplication)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure')
111		29	i	Analog setpoint zero scale	analog setpoint output correction factor zero scale (DSCM-A 0=0 other 32767=0)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure')
112		30	i	Analog setpoint full scale	analog setpoint output correction factor full scale (2000=1*multiplication)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure')
113	115	8	c	Reset	reset facilities (program/alarm/batchcounter)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)')
114	113	6	c	User tag	user definable alias string	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string')
115	97	1	i	Alarm limit maximum	maximum limit for sensor signal to trigger alarm situation	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation')
116	97	2	i	Alarm limit minimum	minimum limit for sensor signal to trigger alarm situation	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation')
117	97	3	c	Alarm mode	alarm mode	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode')
118	97	4	c	Alarm output mode	alarm relais activity mode during alarm situation	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation')
119	97	5	c	Alarm setpoint mode	setpoint change enable during alarm situation	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation')
120	97	6	i	Alarm new setpoint	new/safe setpoint during alarm situation (until reset)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)')
121	104	1	f	Counter value	actual counter value	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value')
122	104	2	c	Counter unit index	pointer to counter unit (classic counter unit table)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)')
123	104	3	f	Counter limit	counter limit/batch	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch')
124	104	4	c	Counter output mode	counter relais activity mode when limit/batch has been reached	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached')
125	104	5	c	Counter setpoint mode	setpoint change enable during counter limit/batch situation (until reset)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)')
126	104	6	i	Counter new setpoint	new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)')
127	104	7	c	Counter unit	counter readout unit (informative only for older devices)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)')
128	1	31	c	Capacity unit	capacity readout unit (informative only for older devices)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)')
129	104	8	c	Counter mode	counter mode	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode')
130	113	7	c	Minimum hardware revision	minimum required hardware revision level for firmware version	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version')
131		1	f	<RCreadfact>	readout factor for direct reading (changes with readunit: local on module, R.O.)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version')
132		2	c	<channumber>	channel number for operation	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version')
133		3	c	<masterchan>	master channel for master-slave operation	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version')
134		4	i	<RCslavefct>	RC slave facto	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version')
135		5	c	<inputnode>	physical node address for channel number	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version')
136		6	c	<inputproc>	physical process for channel number	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version')
137		7	c	<RCreadunit>	readout unit for direct reading (local variable on module: read only)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version')
138	33	1	f	Slave factor	slave factor for master slave control (setp = master output * slave factor)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)')
139	33	2	i	Reference voltage input	reference voltage input for setpoint signal	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal')
140	114	17	c	Stable situation controller response	controller response when controller is stable: |measure-setpoint| < 2%	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%')
141	33	7	f	Temperature	absolute temperature in degrees Celsius	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius')
142	33	8	f	Pressure	absolute pressure in mbar	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar')
143	33	9	f	Time	time in milliseconds	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds')
144	33	10	f	Calibrated volume	calibrated volume in litres	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres')
145		16	c	Sensor number	pointer to sensor number in calibration tube (FPP)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres')
146	115	10	c	Range select	Piston Prover operation mode (write) and status information (read back)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)')
147		2	i	Time out	maximum admitted duration time for specific procedure (in 100 ms)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)')
148	33	9	f	Frequency	frequency in Hz	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz')
149	33	10	f	Impulses/m3	For FRM and FTM imp/m3 and for FCM imp/kg	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg')
150	33	5	f	Normal volume flow	volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min')
151	33	6	f	Volume flow	volume flow at actual conditions in l/min	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min')
152	33	11	f	Delta-p	relative pressure between atmosphere and sensor position	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position')
153	33	13	i	<scalefact>	scaling factor (multiplication) for readout on display (for optimal resolution)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)')
154		17	c	Sensor name	label with information about stop sensor	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)')
155	97	9	c	Reset alarm enable	enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic')
156	104	9	c	Reset counter enable	enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic')
157	33	14	c	Master node	node number of master instrument output signal for a slave	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave')
158	33	15	c	Master process	process number of master instrument output signal for a slave	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave')
159	33	16	c	Remote instrument node	node number of instrument to be operated by another module (keyboard/display)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)')
160	33	17	c	Remote instrument process	process number of instrument to be operated by another module (keyboard/display)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)')
161	33	18	f	Minimum custom range	Mnimum value at 0% for special user readout unit	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit')
162	33	20	f	Maximum custom range	Maximum value at 100% for special user readout unit	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit')
163	115	11	c	Relay/TTL output	Relay/TTL output setting (disabled when used by alarm or counter)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)')
164	114	18	c	Open from zero controller response	Controller response when valve opens from zero	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero')
165	114	20	c	Controller features	Controller settings for special purpose	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose')
166	114	21	f	PID-Kp	PID factor Kp	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp')
167	114	22	f	PID-Ti	PID factor Ti	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti')
168	114	23	f	PID-Td	PID factor Td	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td')
169	33	21	f	Density	Density of selected fluid in kg/m3	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3')
170	113	8	c	Calibration certificate	Number of calibration certificate (last basic calibration)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)')
171	113	9	c	Calibration date	Date of last (basic) calibration	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration')
172	113	10	c	Service number	Service number for repair/rebuilding/recalibration	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration')
173	113	11	c	Service date	Date of last service action	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action')
174	113	12	c	Identification number	Identification number (type) of instrument/device	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device')
175	118	11	c	BHT11	special BHT parameter (to be changed by Bronkhorst only)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)')
176	115	12	c	Power mode	power supply indication in Vdc	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc')
177	113	13	f	Pressure inlet	pressure inlet (upstream) of fluid in bara (for first fluidnr only)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)')
178	113	14	f	Pressure outlet	pressure outlet (downstream) of fluid in bara (for first fluidnr only)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)')
179	113	15	f	Orifice	orifice diameter in mm	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm')
180	113	16	f	Fluid temperature	temperature of fluid through instrument (for first fluidnr only)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)')
181	97	7	c	Alarm delay	time alarm and reset action will be delayed when alarm limit has been exceeded	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded')
182	33	22	f	Capacity 0%	capacity of instrument at zero 0% in sensor base units (mostly equal to zero)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)')
183	0	18	c	Number of channels	number of instrument channels available for this device	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device')
184	0	20	c	Device function	function of device	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device')
185	123	1	c	Scan channel	Channel number to scan with real time information (to be set once)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)')
186	123	3	c	Scan parameter	Parameter number to scan with real time information (to be set once)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)')
187	123	4	i	Scan time	Scan interval time in msec between two samples (to be set once)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)')
188	123	10	c	Scan data	Scanned data with time label (can be readout event by event)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)')
189	114	24	f	Valve open	First-step offset current/voltage for valve when opening from 0%	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%')
190	115	13	c	Number of runs	Amount of runs of a piston prover (0 = stability check)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)')
191	115	14	c	Minimum process time	Minimum process time of a piston prover in 0.1 seconds	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds')
192	116	9	f	Leak rate	Leak rate piston prover	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover')
193	115	15	c	Mode info request	Sets instr. in info mode for 1 read-cycle to check available parameter options	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options')
194	115	16	c	Mode info option list	Gives info about possible values of a mode in an array as result of ModeInfo req	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req')
195	115	17	c	Mode info option description	Gives description about one of the mode options	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options')
196	115	18	c	Calibrations options	Enables/disables options for calibration device (8 bits for 8 options)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)')
197	33	4	f	Mass flow	Real mass flow in kg/min	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min')
198	125	10	c	Bus address	Fieldbus address (top interface)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)')
199	125	3	c	Interface configuration	Configuration setting for interface to other bus-systems	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems')
200	125	9	l	Baudrate	Fieldbus baudrate (top interface)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)')
201	125	20	c	Bus diagnostic string	Fieldbus diagnose string (top interface)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)')
202	115	22	c	Number of vanes	Number of vanes for use in a rotor meter	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter')
203	125	21	c	Fieldbus	Fieldbus name (top interface)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)')
204	33	0	f	fMeasure	measured value for direct reading (in capunits, max.= capacity)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)')
205	33	3	f	fSetpoint	setpoint: wanted value for direct reading (in capunits, max.= capacity)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)')
206	33	23	f	Mass	Mass in g	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g')
207	119	1	c	Manufacturer status register	Manufacturer Status register (64 diagnostic bits)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)')
208	119	2	c	Manufacturer warning register	Manufacturer Warning register (64 diagnostic bits)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)')
209	119	3	c	Manufacturer error register	Manufacturer Error register (64 diagnostic bits)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)')
210	119	4	c	Diagnostic history string	Diagnostic history string (contains history of diag codes)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)')
211	119	5	c	Diagnostic mode	Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)')
212	119	6	c	Manufacturer status enable	Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)')
213	116	21	f	Analog output zero adjust	Analog measure output, zero adjust	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust')
214	116	22	f	Analog output span adjust	Analog measure output, span adjust	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust')
215	116	23	f	Analog input zero adjust	Analog setpoint input, zero adjust	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust')
216	116	24	f	Analog input span adjust	Analog setpoint input, span adjust	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust')
217	116	25	f	Sensor input zero adjust	Sensor input, zero adjust	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust')
218	116	26	f	Sensor input span adjust	Sensor input, span adjust	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust')
219	116	27	f	Temperature input zero adjust	Sensor temperature input, zero adjust	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust')
220	116	28	f	Temperature input span adjust	Sensor temperature input, span adjust	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust')
221	117	5	f	Adaptive smoothing factor	Sensor input filter adapt setting	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting')
222	33	24	i	Slope setpoint step	Slope setpoint step. Setpoint step for the given slopetime	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime')
223	117	6	i	Filter length	Number of samples for average filter	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter')
224	33	25	f	Absolute accuracy	Actual accuracy in current unit	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit')
225	33	26	c	Lookup table index	Lookup table for linearisation index (x and y direction)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)')
226	33	27	f	Lookup table X	Lookup table for linearisation x	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x')
227	33	28	f	Lookup table Y	Lookup table for linearisation y	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y')
228	33	29	c	Lookup table temperature index	Lookup table for linearisation at certain temperature index (z direction)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)')
229	33	30	f	Lookup table temperature	Lookup table for linearisation at certain temperature (z)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)')
230	114	25	f	Valve maximum	Maximum current/voltage for valve	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve')
231	114	26	c	Valve mode	Valve output mode selection	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection')
232	114	27	f	Valve open correction	Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)')
233	114	28	f	Valve zero hold	Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)')
234	114	29	f	Valve slope	Valve slope time (seconds)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)')
235	0	21	c	IFI data	IFI data dump protocol communication string	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string')
236	115	20	c	Range used	Piston Prover information about used sensors	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors')
237	33	31	c	Fluidset properties	Fluidset properties	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties')
238	33	12	c	Lookup table unit type index	Lookup table unit type	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type')
239	33	13	c	Lookup table unit type	Lookup table unit type name	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name')
240	33	16	c	Lookup table unit index	Lookup table unit (unit LUTy)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)')
241	33	17	c	Lookup table unit	Lookup table unit name	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name')
242		29	c	Capacity unit type index	Capacity (readout) unit type	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name')
243		30	c	Capacity unit type	Capacity (readout) unit type name	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name')
244	33	10	f	Capacity unit type temperature	Capacity (readout) unit type temperature (°C)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)')
245	33	11	f	Capacity unit pressure	Capacity (readout) unit type pressure (bar (a))	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))')
246		27	f	Capacity minimum	Minimum capacity in output capacity units	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))')
247		28	f	Capacity maximum	Maximum capacity in output capacity units	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))')
248	113	17	i	Formula type	Formula type needed for conversion	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion')
249	113	18	f	Heat capacity	Heat capacity (Cp) (sensor conditions)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)')
250	113	20	f	Thermal conductivity	Thermal conductivity (sensor conditions)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)')
251	113	21	f	Viscosity	Dynamic viscosity (fluid conditions)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)')
252	113	22	f	Standard flow	Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent')
253	114	30	f	Controller speed	Controller speed factor (gain)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)')
254	113	23	i	Sensor code	Sensor code	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code')
255	113	24	c	Sensor configuration code	Sensor configuration code	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code')
256	113	25	i	Restriction code	Restriction code	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code')
257	113	26	c	Restriction configurator code	Restriction configuration code	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code')
258	113	27	l	Restriction NxP	Restriction NxP (proportional to air equivalent capacity of LFE)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)')
259	113	28	c	Seals information	Seals information (1st byte = other, 2nd = plunger seal)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)')
260	113	29	i	Valve code	Valve code	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code')
261	113	30	c	Valve configuration code	Valve configuration code	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code')
262	113	31	l	Instrument properties	Instrument properties	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties')
263	116	10	c	Lookup table frequency index	Lookup table for frequency index	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index')
264	116	11	f	Lookup table frequency frequency	Lookup table for frequency frequency	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency')
265	116	12	f	Lookup table frequency temperature	Lookup table for frequency temperature	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature')
266	116	13	f	Lookup table frequency density	Lookup table for frequency density	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density')
267	116	14	f	Lookup table frequency span adjust	Lookup table for frequency span adjust	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust')
268	65	15	c	Capacity unit index (ext)	Capacity (readout) unit index (extended unit table)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)')
269	116	15	f	Density actual	Actual density, measured by instrument	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument')
270	116	18	f	Measured restriction	Measured restriction	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction')
271	116	8	c	Temperature potmeter	Potmeter for sensor temperature compensation	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation')
272	116	9	c	Temperature potmeter gain	Gain for sensor temperature compensation	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation')
273	104	10	f	Counter controller overrun correction	Counter controller overrun correction	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction')
274	104	11	f	Counter controller gain	Counter controller gain	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain')
275	65	1	c	Sub fluid number	Sub fluid number	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number')
276	116	17	f	Temperature compensation factor	Temperature compensation factor	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor')
277	116	29	l	DSP register address	DSP register address	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address')
278	116	30	l	DSP register long	DSP register long	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long')
279	116	30	f	DSP register floating point	DSP register floating point	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point')
280	116	31	i	DSP register integer	DSP register integer	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer')
281	121	0	f	Standard deviation	Standard deviation	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation')
282	121	1	i	Measurement status	Measurement status	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status')
283	121	2	i	Measurement stop criteria	Measurement stop criteria	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria')
284	121	3	i	Measurement time out	Measurement time out	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out')
285	121	4	i	Maximum number of runs	Measurement maximum number of runs	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs')
286	121	5	f	Minimum standard deviation	Measurement minimum standard deviation	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation')
287	114	31	l	IO switch status	IO status for switches	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches')
288	65	21	i	Sensor bridge settings	Sensor bridge settings	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings')
289	65	22	f	Sensor bridge current	Sensor bridge current (0..0.02A)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)')
290	65	23	f	Sensor resistance	Sensor resistance (Ohm)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)')
291	65	24	f	Sensor bridge voltage	Sensor bridge voltage (0..11.5V)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)')
292	65	25	c	Sensor group name	Sensor name (based on sensor detection / sensor resistance)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)')
293	116	20	f	Sensor calibration temperature	Sensor temperature at calibration	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration')
294	115	31	c	Valve safe state	Valve safe state	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state')
295	104	12	c	Counter unit type index	Counter unit type	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type')
296	104	13	c	Counter unit type	Counter unit type name	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name')
297	104	14	c	Counter unit index (ext)	Counter unit index (extended counter unit table)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)')
298	125	8	c	Bus1 selection	Fieldbus select	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select')
299	125	11	c	Bus1 medium	Fieldbus medium id	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id')
300	124	7	c	Bus2 mode	Fieldbus2 mode	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode')
301	124	8	c	Bus2 selection	Fieldbus2 select	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select')
302	124	10	c	Bus2 address	Fieldbus2 address (side interface)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)')
303	124	9	l	Bus2 baudrate	Fieldbus2 baudrate (side interface)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)')
304	124	11	c	Bus2 medium	Fieldbus2 medium id	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id')
305	124	20	c	Bus2 diagnostics	Fieldbus2 diagnose string (side interface)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)')
306	124	21	c	Bus2 name	Fieldbus2 name (side interface)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)')
307	120	0	c	PIO channel selection	PIO channel selection	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection')
308	120	2	i	PIO parameter	PIO parameter connected to i/o channel (=process number*256+parameter number)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)')
309	120	6	f	PIO input/output filter	PIO input/output filter constant (0..1, 0=max, 1=min)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)')
310	120	7	f	PIO parameter capacity 0%	PIO parameter capacity, 0% value	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value')
311	120	3	f	PIO parameter capacity 100%	PIO parameter capacity, 100% value	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value')
312	120	1	c	PIO configuration selection	PIO i/o channel configuration selection	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection')