313	120	4	f	PIO analog zero adjust	PIO analog input/output, zero adjust value	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value')
314	120	5	f	PIO analog span adjust	PIO analog input/output, span adjust factor	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor')
315	120	8	f	PIO hardware capacity max	PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)')
316	120	9	c	PIO capacity set selection	PIO capacity set selection	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection')
317	120	10	f	PIO hardware capacity 0%	PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c'), (317, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity 0%')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)')
318	120	11	f	PIO hardware capacity 100%	PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c'), (317, 'f'), (318, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity 0%'), (318, 'PIO hardware capacity 100%')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)'), (318, 'PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)')
319	0	6	i	Hardware platform id	Hardware platform identification number	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c'), (317, 'f'), (318, 'f'), (319, 'i')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity 0%'), (318, 'PIO hardware capacity 100%'), (319, 'Hardware platform id')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)'), (318, 'PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)'), (319, 'Hardware platform identification number')
320	0	7	c	Hardware platform sub id	Hardware platform sub identification number	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c'), (317, 'f'), (318, 'f'), (319, 'i'), (320, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity 0%'), (318, 'PIO hardware capacity 100%'), (319, 'Hardware platform id'), (320, 'Hardware platform sub id')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)'), (318, 'PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)'), (319, 'Hardware platform identification number'), (320, 'Hardware platform sub identification number')
321	124	31	l	Temporary baudrate	Temporary volatile Fieldbus2 baudrate (side interface)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c'), (317, 'f'), (318, 'f'), (319, 'i'), (320, 'c'), (321, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity 0%'), (318, 'PIO hardware capacity 100%'), (319, 'Hardware platform id'), (320, 'Hardware platform sub id'), (321, 'Temporary baudrate')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)'), (318, 'PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)'), (319, 'Hardware platform identification number'), (320, 'Hardware platform sub identification number'), (321, 'Temporary volatile Fieldbus2 baudrate (side interface)')
322	115	23	c	Setpoint monitor mode	Monitor mode for setpoint	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c'), (317, 'f'), (318, 'f'), (319, 'i'), (320, 'c'), (321, 'l'), (322, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity 0%'), (318, 'PIO hardware capacity 100%'), (319, 'Hardware platform id'), (320, 'Hardware platform sub id'), (321, 'Temporary baudrate'), (322, 'Setpoint monitor mode')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)'), (318, 'PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)'), (319, 'Hardware platform identification number'), (320, 'Hardware platform sub identification number'), (321, 'Temporary volatile Fieldbus2 baudrate (side interface)'), (322, 'Monitor mode for setpoint')
323	118	12	l	BHT12	Special BHT parameter	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c'), (317, 'f'), (318, 'f'), (319, 'i'), (320, 'c'), (321, 'l'), (322, 'c'), (323, 'l')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity 0%'), (318, 'PIO hardware capacity 100%'), (319, 'Hardware platform id'), (320, 'Hardware platform sub id'), (321, 'Temporary baudrate'), (322, 'Setpoint monitor mode'), (323, 'BHT12')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)'), (318, 'PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)'), (319, 'Hardware platform identification number'), (320, 'Hardware platform sub identification number'), (321, 'Temporary volatile Fieldbus2 baudrate (side interface)'), (322, 'Monitor mode for setpoint'), (323, 'Special BHT parameter')
324	65	26	f	Nominal sensor voltage	Nominal sensor voltage (used for sensor temperature voltage compensation)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c'), (317, 'f'), (318, 'f'), (319, 'i'), (320, 'c'), (321, 'l'), (322, 'c'), (323, 'l'), (324, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity 0%'), (318, 'PIO hardware capacity 100%'), (319, 'Hardware platform id'), (320, 'Hardware platform sub id'), (321, 'Temporary baudrate'), (322, 'Setpoint monitor mode'), (323, 'BHT12'), (324, 'Nominal sensor voltage')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)'), (318, 'PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)'), (319, 'Hardware platform identification number'), (320, 'Hardware platform sub identification number'), (321, 'Temporary volatile Fieldbus2 baudrate (side interface)'), (322, 'Monitor mode for setpoint'), (323, 'Special BHT parameter'), (324, 'Nominal sensor voltage (used for sensor temperature voltage compensation)')
325	116	16	f	Sensor voltage compensation factor	Sensor voltage compensation factor (used for sensor temperature voltage comp.)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c'), (317, 'f'), (318, 'f'), (319, 'i'), (320, 'c'), (321, 'l'), (322, 'c'), (323, 'l'), (324, 'f'), (325, 'f')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity 0%'), (318, 'PIO hardware capacity 100%'), (319, 'Hardware platform id'), (320, 'Hardware platform sub id'), (321, 'Temporary baudrate'), (322, 'Setpoint monitor mode'), (323, 'BHT12'), (324, 'Nominal sensor voltage'), (325, 'Sensor voltage compensation factor')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)'), (318, 'PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)'), (319, 'Hardware platform identification number'), (320, 'Hardware platform sub identification number'), (321, 'Temporary volatile Fieldbus2 baudrate (side interface)'), (322, 'Monitor mode for setpoint'), (323, 'Special BHT parameter'), (324, 'Nominal sensor voltage (used for sensor temperature voltage compensation)'), (325, 'Sensor voltage compensation factor (used for sensor temperature voltage comp.)')
326	119	31	c	PCB serial number	PCB serial number (unique number or string written by PCB manufacturer)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c'), (317, 'f'), (318, 'f'), (319, 'i'), (320, 'c'), (321, 'l'), (322, 'c'), (323, 'l'), (324, 'f'), (325, 'f'), (326, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity 0%'), (318, 'PIO hardware capacity 100%'), (319, 'Hardware platform id'), (320, 'Hardware platform sub id'), (321, 'Temporary baudrate'), (322, 'Setpoint monitor mode'), (323, 'BHT12'), (324, 'Nominal sensor voltage'), (325, 'Sensor voltage compensation factor'), (326, 'PCB serial number')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)'), (318, 'PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)'), (319, 'Hardware platform identification number'), (320, 'Hardware platform sub identification number'), (321, 'Temporary volatile Fieldbus2 baudrate (side interface)'), (322, 'Monitor mode for setpoint'), (323, 'Special BHT parameter'), (324, 'Nominal sensor voltage (used for sensor temperature voltage compensation)'), (325, 'Sensor voltage compensation factor (used for sensor temperature voltage comp.)'), (326, 'PCB serial number (unique number or string written by PCB manufacturer)')
327	115	24	c	Minimum measure time	Minimum measure time between sensors of a piston prover in 0.1 seconds	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c'), (317, 'f'), (318, 'f'), (319, 'i'), (320, 'c'), (321, 'l'), (322, 'c'), (323, 'l'), (324, 'f'), (325, 'f'), (326, 'c'), (327, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity 0%'), (318, 'PIO hardware capacity 100%'), (319, 'Hardware platform id'), (320, 'Hardware platform sub id'), (321, 'Temporary baudrate'), (322, 'Setpoint monitor mode'), (323, 'BHT12'), (324, 'Nominal sensor voltage'), (325, 'Sensor voltage compensation factor'), (326, 'PCB serial number'), (327, 'Minimum measure time')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)'), (318, 'PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)'), (319, 'Hardware platform identification number'), (320, 'Hardware platform sub identification number'), (321, 'Temporary volatile Fieldbus2 baudrate (side interface)'), (322, 'Monitor mode for setpoint'), (323, 'Special BHT parameter'), (324, 'Nominal sensor voltage (used for sensor temperature voltage compensation)'), (325, 'Sensor voltage compensation factor (used for sensor temperature voltage comp.)'), (326, 'PCB serial number (unique number or string written by PCB manufacturer)'), (327, 'Minimum measure time between sensors of a piston prover in 0.1 seconds')
328	125	12	c	Bus1 parity	Fieldbus parity	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c'), (317, 'f'), (318, 'f'), (319, 'i'), (320, 'c'), (321, 'l'), (322, 'c'), (323, 'l'), (324, 'f'), (325, 'f'), (326, 'c'), (327, 'c'), (328, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity 0%'), (318, 'PIO hardware capacity 100%'), (319, 'Hardware platform id'), (320, 'Hardware platform sub id'), (321, 'Temporary baudrate'), (322, 'Setpoint monitor mode'), (323, 'BHT12'), (324, 'Nominal sensor voltage'), (325, 'Sensor voltage compensation factor'), (326, 'PCB serial number'), (327, 'Minimum measure time'), (328, 'Bus1 parity')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)'), (318, 'PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)'), (319, 'Hardware platform identification number'), (320, 'Hardware platform sub identification number'), (321, 'Temporary volatile Fieldbus2 baudrate (side interface)'), (322, 'Monitor mode for setpoint'), (323, 'Special BHT parameter'), (324, 'Nominal sensor voltage (used for sensor temperature voltage compensation)'), (325, 'Sensor voltage compensation factor (used for sensor temperature voltage comp.)'), (326, 'PCB serial number (unique number or string written by PCB manufacturer)'), (327, 'Minimum measure time between sensors of a piston prover in 0.1 seconds'), (328, 'Fieldbus parity')
329	124	12	c	Bus2 parity	Fieldbus2 parity (side interface)	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c'), (317, 'f'), (318, 'f'), (319, 'i'), (320, 'c'), (321, 'l'), (322, 'c'), (323, 'l'), (324, 'f'), (325, 'f'), (326, 'c'), (327, 'c'), (328, 'c'), (329, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity 0%'), (318, 'PIO hardware capacity 100%'), (319, 'Hardware platform id'), (320, 'Hardware platform sub id'), (321, 'Temporary baudrate'), (322, 'Setpoint monitor mode'), (323, 'BHT12'), (324, 'Nominal sensor voltage'), (325, 'Sensor voltage compensation factor'), (326, 'PCB serial number'), (327, 'Minimum measure time'), (328, 'Bus1 parity'), (329, 'Bus2 parity')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)'), (318, 'PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)'), (319, 'Hardware platform identification number'), (320, 'Hardware platform sub identification number'), (321, 'Temporary volatile Fieldbus2 baudrate (side interface)'), (322, 'Monitor mode for setpoint'), (323, 'Special BHT parameter'), (324, 'Nominal sensor voltage (used for sensor temperature voltage compensation)'), (325, 'Sensor voltage compensation factor (used for sensor temperature voltage comp.)'), (326, 'PCB serial number (unique number or string written by PCB manufacturer)'), (327, 'Minimum measure time between sensors of a piston prover in 0.1 seconds'), (328, 'Fieldbus parity'), (329, 'Fieldbus2 parity (side interface)')
330	0	8	c	Firmware id	Firmware identification number	Var, (0, 'c'), (1, 'c'), (2, 'c'), (3, 'c'), (4, 'c'), (5, 'c'), (6, 'c'), (7, 'i'), (8, 'i'), (9, 'i'), (10, 'i'), (11, 'c'), (12, 'f'), (13, 'f'), (14, 'f'), (15, 'f'), (16, 'f'), (17, 'f'), (18, 'f'), (19, 'f'), (20, 'f'), (21, 'c'), (22, 'c'), (23, 'c'), (24, 'c'), (25, 'c'), (26, 'c'), (27, 'c'), (28, 'c'), (29, 'c'), (30, 'c'), (31, 'c'), (32, 'c'), (33, 'c'), (34, 'c'), (51, 'c'), (52, 'c'), (53, 'i'), (54, 'l'), (55, 'f'), (56, 'f'), (57, 'c'), (58, 'i'), (59, 'c'), (60, 'c'), (61, 'c'), (62, 'l'), (63, 'l'), (64, 'l'), (65, 'l'), (66, 'l'), (67, 'c'), (68, 'c'), (69, 'c'), (70, 'c'), (71, 'c'), (72, 'f'), (73, 'f'), (78, 'c'), (79, 'c'), (80, 'c'), (81, 'f'), (82, 'c'), (83, 'c'), (84, 'c'), (85, 'c'), (86, 'c'), (87, 'c'), (88, 'c'), (89, 'c'), (90, 'c'), (91, 'c'), (92, 'c'), (93, 'c'), (94, 'i'), (95, 'l'), (96, 'i'), (97, 'c'), (98, 'c'), (99, 'c'), (100, 'c'), (101, 'l'), (102, 'c'), (103, 'c'), (104, 'c'), (105, 'c'), (106, 'f'), (113, 'c'), (114, 'c'), (115, 'i'), (116, 'i'), (117, 'c'), (118, 'c'), (119, 'c'), (120, 'i'), (121, 'f'), (122, 'c'), (123, 'f'), (124, 'c'), (125, 'c'), (126, 'i'), (127, 'c'), (128, 'c'), (129, 'c'), (130, 'c'), (138, 'f'), (139, 'i'), (140, 'c'), (141, 'f'), (142, 'f'), (143, 'f'), (144, 'f'), (146, 'c'), (148, 'f'), (149, 'f'), (150, 'f'), (151, 'f'), (152, 'f'), (153, 'i'), (155, 'c'), (156, 'c'), (157, 'c'), (158, 'c'), (159, 'c'), (160, 'c'), (161, 'f'), (162, 'f'), (163, 'c'), (164, 'c'), (165, 'c'), (166, 'f'), (167, 'f'), (168, 'f'), (169, 'f'), (170, 'c'), (171, 'c'), (172, 'c'), (173, 'c'), (174, 'c'), (175, 'c'), (176, 'c'), (177, 'f'), (178, 'f'), (179, 'f'), (180, 'f'), (181, 'c'), (182, 'f'), (183, 'c'), (184, 'c'), (185, 'c'), (186, 'c'), (187, 'i'), (188, 'c'), (189, 'f'), (190, 'c'), (191, 'c'), (192, 'f'), (193, 'c'), (194, 'c'), (195, 'c'), (196, 'c'), (197, 'f'), (198, 'c'), (199, 'c'), (200, 'l'), (201, 'c'), (202, 'c'), (203, 'c'), (204, 'f'), (205, 'f'), (206, 'f'), (207, 'c'), (208, 'c'), (209, 'c'), (210, 'c'), (211, 'c'), (212, 'c'), (213, 'f'), (214, 'f'), (215, 'f'), (216, 'f'), (217, 'f'), (218, 'f'), (219, 'f'), (220, 'f'), (221, 'f'), (222, 'i'), (223, 'i'), (224, 'f'), (225, 'c'), (226, 'f'), (227, 'f'), (228, 'c'), (229, 'f'), (230, 'f'), (231, 'c'), (232, 'f'), (233, 'f'), (234, 'f'), (235, 'c'), (236, 'c'), (237, 'c'), (238, 'c'), (239, 'c'), (240, 'c'), (241, 'c'), (244, 'f'), (245, 'f'), (248, 'i'), (249, 'f'), (250, 'f'), (251, 'f'), (252, 'f'), (253, 'f'), (254, 'i'), (255, 'c'), (256, 'i'), (257, 'c'), (258, 'l'), (259, 'c'), (260, 'i'), (261, 'c'), (262, 'l'), (263, 'c'), (264, 'f'), (265, 'f'), (266, 'f'), (267, 'f'), (268, 'c'), (269, 'f'), (270, 'f'), (271, 'c'), (272, 'c'), (273, 'f'), (274, 'f'), (275, 'c'), (276, 'f'), (277, 'l'), (278, 'l'), (279, 'f'), (280, 'i'), (281, 'f'), (282, 'i'), (283, 'i'), (284, 'i'), (285, 'i'), (286, 'f'), (287, 'l'), (288, 'i'), (289, 'f'), (290, 'f'), (291, 'f'), (292, 'c'), (293, 'f'), (294, 'c'), (295, 'c'), (296, 'c'), (297, 'c'), (298, 'c'), (299, 'c'), (300, 'c'), (301, 'c'), (302, 'c'), (303, 'l'), (304, 'c'), (305, 'c'), (306, 'c'), (307, 'c'), (308, 'i'), (309, 'f'), (310, 'f'), (311, 'f'), (312, 'c'), (313, 'f'), (314, 'f'), (315, 'f'), (316, 'c'), (317, 'f'), (318, 'f'), (319, 'i'), (320, 'c'), (321, 'l'), (322, 'c'), (323, 'l'), (324, 'f'), (325, 'f'), (326, 'c'), (327, 'c'), (328, 'c'), (329, 'c'), (330, 'c')	Name, (0, 'Identification string'), (1, 'Primary node address'), (2, 'Secondary node address'), (3, 'Next node address'), (4, 'Last node address'), (5, 'Arbitrage'), (6, 'Initreset'), (7, 'Measure'), (8, 'Setpoint'), (9, 'Setpoint slope'), (10, 'Analog input'), (11, 'Control mode'), (12, 'Polynomial constant A'), (13, 'Polynomial constant B'), (14, 'Polynomial constant C'), (15, 'Polynomial constant D'), (16, 'Polynomial constant E'), (17, 'Polynomial constant F'), (18, 'Polynomial constant G'), (19, 'Polynomial constant H'), (20, 'Capacity'), (21, 'Sensor type'), (22, 'Capacity unit index'), (23, 'Fluid number'), (24, 'Fluid name'), (25, 'Claim node'), (26, 'Modify'), (27, 'Alarm info'), (28, 'Channel amount'), (29, 'First channel'), (30, 'Last channel'), (31, '<hostcontrl>'), (32, 'Alarm message unit type'), (33, 'Alarm message number'), (34, 'Relay status'), (51, 'Cycle time'), (52, 'Analog mode'), (53, 'Reference voltage'), (54, 'Valve output'), (55, 'Dynamic display factor'), (56, 'Static display factor'), (57, 'Calibration mode'), (58, 'Valve offset'), (59, 'Monitor mode'), (60, 'Alarm register1'), (61, 'Alarm register2'), (62, '<CalRegZS1>'), (63, '<CalRegFS1>'), (64, '<CalRegZS2>'), (65, '<CalRegFS2>'), (66, 'ADC control register'), (67, 'Bridge potmeter'), (68, '<AlarmEnble>'), (69, 'Test mode'), (70, '<ADC channel select>'), (71, 'Normal step controller response'), (72, 'Setpoint exponential smoothing filter'), (73, 'Sensor exponential smoothing filter'), (78, 'Tuning mode'), (79, 'Valve default'), (80, 'Global modify'), (81, 'Valve span correction factor'), (82, 'Valve curve correction'), (83, '<MemShipNor>'), (84, '<MemShipOpn>'), (85, 'IO status'), (86, '<FuzzStNeNo>'), (87, '<FuzzStPoNo>'), (88, '<FuzzStOpen>'), (89, 'Device type'), (90, 'BHTModel number'), (91, 'Serial number'), (92, 'Customer model'), (93, 'BHT1'), (94, 'BHT2'), (95, 'BHT3'), (96, 'BHT4'), (97, 'BHT5'), (98, 'BHT6'), (99, 'BHT7'), (100, 'BHT8'), (101, 'BHT9'), (102, 'BHT10'), (103, 'Broadcast repeating time'), (104, 'Firmware version'), (105, 'Pressure sensor type'), (106, 'Barometer pressure'), (113, 'Reset'), (114, 'User tag'), (115, 'Alarm limit maximum'), (116, 'Alarm limit minimum'), (117, 'Alarm mode'), (118, 'Alarm output mode'), (119, 'Alarm setpoint mode'), (120, 'Alarm new setpoint'), (121, 'Counter value'), (122, 'Counter unit index'), (123, 'Counter limit'), (124, 'Counter output mode'), (125, 'Counter setpoint mode'), (126, 'Counter new setpoint'), (127, 'Counter unit'), (128, 'Capacity unit'), (129, 'Counter mode'), (130, 'Minimum hardware revision'), (138, 'Slave factor'), (139, 'Reference voltage input'), (140, 'Stable situation controller response'), (141, 'Temperature'), (142, 'Pressure'), (143, 'Time'), (144, 'Calibrated volume'), (146, 'Range select'), (148, 'Frequency'), (149, 'Impulses/m3'), (150, 'Normal volume flow'), (151, 'Volume flow'), (152, 'Delta-p'), (153, '<scalefact>'), (155, 'Reset alarm enable'), (156, 'Reset counter enable'), (157, 'Master node'), (158, 'Master process'), (159, 'Remote instrument node'), (160, 'Remote instrument process'), (161, 'Minimum custom range'), (162, 'Maximum custom range'), (163, 'Relay/TTL output'), (164, 'Open from zero controller response'), (165, 'Controller features'), (166, 'PID-Kp'), (167, 'PID-Ti'), (168, 'PID-Td'), (169, 'Density'), (170, 'Calibration certificate'), (171, 'Calibration date'), (172, 'Service number'), (173, 'Service date'), (174, 'Identification number'), (175, 'BHT11'), (176, 'Power mode'), (177, 'Pressure inlet'), (178, 'Pressure outlet'), (179, 'Orifice'), (180, 'Fluid temperature'), (181, 'Alarm delay'), (182, 'Capacity 0%'), (183, 'Number of channels'), (184, 'Device function'), (185, 'Scan channel'), (186, 'Scan parameter'), (187, 'Scan time'), (188, 'Scan data'), (189, 'Valve open'), (190, 'Number of runs'), (191, 'Minimum process time'), (192, 'Leak rate'), (193, 'Mode info request'), (194, 'Mode info option list'), (195, 'Mode info option description'), (196, 'Calibrations options'), (197, 'Mass flow'), (198, 'Bus address'), (199, 'Interface configuration'), (200, 'Baudrate'), (201, 'Bus diagnostic string'), (202, 'Number of vanes'), (203, 'Fieldbus'), (204, 'fMeasure'), (205, 'fSetpoint'), (206, 'Mass'), (207, 'Manufacturer status register'), (208, 'Manufacturer warning register'), (209, 'Manufacturer error register'), (210, 'Diagnostic history string'), (211, 'Diagnostic mode'), (212, 'Manufacturer status enable'), (213, 'Analog output zero adjust'), (214, 'Analog output span adjust'), (215, 'Analog input zero adjust'), (216, 'Analog input span adjust'), (217, 'Sensor input zero adjust'), (218, 'Sensor input span adjust'), (219, 'Temperature input zero adjust'), (220, 'Temperature input span adjust'), (221, 'Adaptive smoothing factor'), (222, 'Slope setpoint step'), (223, 'Filter length'), (224, 'Absolute accuracy'), (225, 'Lookup table index'), (226, 'Lookup table X'), (227, 'Lookup table Y'), (228, 'Lookup table temperature index'), (229, 'Lookup table temperature'), (230, 'Valve maximum'), (231, 'Valve mode'), (232, 'Valve open correction'), (233, 'Valve zero hold'), (234, 'Valve slope'), (235, 'IFI data'), (236, 'Range used'), (237, 'Fluidset properties'), (238, 'Lookup table unit type index'), (239, 'Lookup table unit type'), (240, 'Lookup table unit index'), (241, 'Lookup table unit'), (244, 'Capacity unit type temperature'), (245, 'Capacity unit pressure'), (248, 'Formula type'), (249, 'Heat capacity'), (250, 'Thermal conductivity'), (251, 'Viscosity'), (252, 'Standard flow'), (253, 'Controller speed'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configurator code'), (258, 'Restriction NxP'), (259, 'Seals information'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table frequency index'), (264, 'Lookup table frequency frequency'), (265, 'Lookup table frequency temperature'), (266, 'Lookup table frequency density'), (267, 'Lookup table frequency span adjust'), (268, 'Capacity unit index (ext)'), (269, 'Density actual'), (270, 'Measured restriction'), (271, 'Temperature potmeter'), (272, 'Temperature potmeter gain'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Maximum number of runs'), (286, 'Minimum standard deviation'), (287, 'IO switch status'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current'), (290, 'Sensor resistance'), (291, 'Sensor bridge voltage'), (292, 'Sensor group name'), (293, 'Sensor calibration temperature'), (294, 'Valve safe state'), (295, 'Counter unit type index'), (296, 'Counter unit type'), (297, 'Counter unit index (ext)'), (298, 'Bus1 selection'), (299, 'Bus1 medium'), (300, 'Bus2 mode'), (301, 'Bus2 selection'), (302, 'Bus2 address'), (303, 'Bus2 baudrate'), (304, 'Bus2 medium'), (305, 'Bus2 diagnostics'), (306, 'Bus2 name'), (307, 'PIO channel selection'), (308, 'PIO parameter'), (309, 'PIO input/output filter'), (310, 'PIO parameter capacity 0%'), (311, 'PIO parameter capacity 100%'), (312, 'PIO configuration selection'), (313, 'PIO analog zero adjust'), (314, 'PIO analog span adjust'), (315, 'PIO hardware capacity max'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity 0%'), (318, 'PIO hardware capacity 100%'), (319, 'Hardware platform id'), (320, 'Hardware platform sub id'), (321, 'Temporary baudrate'), (322, 'Setpoint monitor mode'), (323, 'BHT12'), (324, 'Nominal sensor voltage'), (325, 'Sensor voltage compensation factor'), (326, 'PCB serial number'), (327, 'Minimum measure time'), (328, 'Bus1 parity'), (329, 'Bus2 parity'), (330, 'Firmware id')	descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)'), (318, 'PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)'), (319, 'Hardware platform identification number'), (320, 'Hardware platform sub identification number'), (321, 'Temporary volatile Fieldbus2 baudrate (side interface)'), (322, 'Monitor mode for setpoint'), (323, 'Special BHT parameter'), (324, 'Nominal sensor voltage (used for sensor temperature voltage compensation)'), (325, 'Sensor voltage compensation factor (used for sensor temperature voltage comp.)'), (326, 'PCB serial number (unique number or string written by PCB manufacturer)'), (327, 'Minimum measure time between sensors of a piston prover in 0.1 seconds'), (328, 'Fieldbus parity'), (329, 'Fieldbus2 parity (side interface)'), (330, 'Firmware identification number')
				dict(descr, (0, 'identnr.+softwareversion[+serialnr.]'), (1, 'primary node address: network parameter FLOW-BUS'), (2, 'secundary node address: network parameter FLOW-BUS'), (3, 'next node address: network parameter FLOW-BUS'), (4, 'last node address: network parameter FLOW-BUS'), (5, 'FLOW-BUS arbitrage setting and/or automatic optimization'), (6, 'init and reset security key commands for network/parameter settings'), (7, 'measured value (100% = 32000)'), (8, 'setpoint: wanted value (100% = 32000)'), (9, 'setpoint ramp signal 0..100 % in up to slope x 0.1 sec.'), (10, 'analog input signal, normally used for ext. setp. (100% = 32000)'), (11, 'control mode selection for instrument or module'), (12, 'polynomial constant A (offset)'), (13, 'polynomial constant B (span]'), (14, 'polynomial constant C'), (15, 'polynomial constant D'), (16, 'polynomial constant E (offset) for setpoint or power value'), (17, 'polynomial constant F (span) for setpoint or power value'), (18, 'polynomial constant G for setpoint or power value'), (19, 'polynomial constant H for setpoint or power value'), (20, 'readout value at 100% in capacity (readout) unit'), (21, 'sensor type information for actual reading and sensor/controller indication'), (22, 'pointer to capacity (readout) unit (classic unit table]'), (23, 'pointer to capacity (readout) unit (classic unit table], c'), (24, 'name of fluid'), (25, 'node address of module with operation rights'), (26, 'contains number(s) of changed par (0xXX par nr, 0xFF more than one par changed)'), (27, 'status information of several alarms/errors in the instrument'), (28, 'amount of channels which can be operated'), (29, 'first channel that can be operated'), (30, 'last channel that can be operated'), (31, 'operation by HOST computer enable flag'), (32, 'alarm message string with unit type information'), (33, 'alarm message string with unit number information'), (34, 'status of relays/potential free contacts ('0' = not activated, '1' = activated)'), (51, 'cycle time * 10 msec. main loop signal processing'), (52, 'analog mode selection for analog operation'), (53, 'reference voltage output signal for analog operation'), (54, 'valve output signal (24-bit number in range 0..14.3Vdc/0..23.3Vdc)'), (55, 'dynamic display factor for display filter (0=max, 1=min goes with par 57)'), (56, 'static display factor for display filter (0=max, 1=min goes with par 56)'), (57, 'calibration mode selection (not active until cntrlmode has been set to value 9)'), (58, 'valve offset: amount of DAC steps within 1 potmeter step'), (59, 'monitor: output signal (measure) selection for bus and analog output'), (60, 'alarm register containing warning flags'), (61, 'alarm register containing critical error flags'), (62, 'calibration register zero scale input 1 ADC'), (63, 'calibration register full scale input 1 ADC'), (64, 'calibration register zero scale input 2 ADC'), (65, 'calibration register full scale input 2 ADC'), (66, 'ADC control register'), (67, 'sensor bridge zero potmeter setting'), (68, 'broadcast alarm messsage enable flag'), (69, 'test mode selection (not active until cntrlmode has been set to value 5)'), (70, 'channel selection ADC'), (71, 'controller response for normal steps (128=normal, <128=slower, >128=faster)'), (72, 'analog input filter constant (0=max, 1=min)'), (73, 'sensor input filter constant (0=max, 1=min)'), (78, '(auto)tuning mode selection (not active until cntrlmode has been set to value 6)'), (79, 'valve type (needed for controlling behaviour)'), (80, 'contains number(s) of changed processes for indirect polling (0xXX / 0xFF)'), (81, 'correction factor valve curve ratio high/low area'), (82, 'Valve curve correction for controller (max. factor*0.1, flow where factor = 1)'), (83, 'array with memberships for normal Fuzzy controller'), (84, 'array with memberships for 0-open Fuzzy controller'), (85, 'IO status byte for jumper settings and LED signal modes'), (86, 'array with neg nor output steps for Fuzzy contr.'), (87, 'array with pos nor output steps for Fuzzy contr.'), (88, 'array with open at 0 output steps for Fuzzy contr.'), (89, '(FLOW-BUS) device type information string'), (90, 'model number information string'), (91, 'serial number information string (to be changed by Bronkhorst only)'), (92, 'manufacturing configuration information string'), (93, 'special BHT parameter (to be changed by Bronkhorst only)'), (94, 'special BHT parameter'), (95, 'special BHT parameter'), (96, 'special BHT parameter'), (97, 'special BHT parameter'), (98, 'special BHT parameter'), (99, 'special BHT parameter'), (100, 'special BHT parameter'), (101, 'special BHT parameter'), (102, 'special BHT parameter (to be changed by Bronkhorst only)'), (103, 'broadcast repeating time (x5 ms) (old:height of open at 0 pulse train for valve)'), (104, 'revision number of firmware'), (105, 'type of pressure sensor'), (106, 'mbar atmospheric (central) barometer pressure'), (113, 'reset facilities (program/alarm/batchcounter)'), (114, 'user definable alias string'), (115, 'maximum limit for sensor signal to trigger alarm situation'), (116, 'minimum limit for sensor signal to trigger alarm situation'), (117, 'alarm mode'), (118, 'alarm relais activity mode during alarm situation'), (119, 'setpoint change enable during alarm situation'), (120, 'new/safe setpoint during alarm situation (until reset)'), (121, 'actual counter value'), (122, 'pointer to counter unit (classic counter unit table)'), (123, 'counter limit/batch'), (124, 'counter relais activity mode when limit/batch has been reached'), (125, 'setpoint change enable during counter limit/batch situation (until reset)'), (126, 'new/safe setpoint at counter limit/batch situation (until reset) (normally = 0%)'), (127, 'counter readout unit (informative only for older devices)'), (128, 'capacity readout unit (informative only for older devices)'), (129, 'counter mode'), (130, 'minimum required hardware revision level for firmware version'), (138, 'slave factor for master slave control (setp = master output * slave factor)'), (139, 'reference voltage input for setpoint signal'), (140, 'controller response when controller is stable: |measure-setpoint| < 2%'), (141, 'absolute temperature in degrees Celsius'), (142, 'absolute pressure in mbar'), (143, 'time in milliseconds'), (144, 'calibrated volume in litres'), (146, 'Piston Prover operation mode (write) and status information (read back)'), (148, 'frequency in Hz'), (149, 'For FRM and FTM imp/m3 and for FCM imp/kg'), (150, 'volume flow referenced to normal conditions i.e. 0 °C, 1013.25 hPa(a) in ln/min'), (151, 'volume flow at actual conditions in l/min'), (152, 'relative pressure between atmosphere and sensor position'), (153, 'scaling factor (multiplication) for readout on display (for optimal resolution)'), (155, 'enable reset of alarm by: keyboard, external signal, FLOW-BUS, automatic'), (156, 'enable reset of counter by: keyboard, external signal, FLOW-BUS, automatic'), (157, 'node number of master instrument output signal for a slave'), (158, 'process number of master instrument output signal for a slave'), (159, 'node number of instrument to be operated by another module (keyboard/display)'), (160, 'process number of instrument to be operated by another module (keyboard/display)'), (161, 'Mnimum value at 0% for special user readout unit'), (162, 'Maximum value at 100% for special user readout unit'), (163, 'Relay/TTL output setting (disabled when used by alarm or counter)'), (164, 'Controller response when valve opens from zero'), (165, 'Controller settings for special purpose'), (166, 'PID factor Kp'), (167, 'PID factor Ti'), (168, 'PID factor Td'), (169, 'Density of selected fluid in kg/m3'), (170, 'Number of calibration certificate (last basic calibration)'), (171, 'Date of last (basic) calibration'), (172, 'Service number for repair/rebuilding/recalibration'), (173, 'Date of last service action'), (174, 'Identification number (type) of instrument/device'), (175, 'special BHT parameter (to be changed by Bronkhorst only)'), (176, 'power supply indication in Vdc'), (177, 'pressure inlet (upstream) of fluid in bara (for first fluidnr only)'), (178, 'pressure outlet (downstream) of fluid in bara (for first fluidnr only)'), (179, 'orifice diameter in mm'), (180, 'temperature of fluid through instrument (for first fluidnr only)'), (181, 'time alarm and reset action will be delayed when alarm limit has been exceeded'), (182, 'capacity of instrument at zero 0% in sensor base units (mostly equal to zero)'), (183, 'number of instrument channels available for this device'), (184, 'function of device'), (185, 'Channel number to scan with real time information (to be set once)'), (186, 'Parameter number to scan with real time information (to be set once)'), (187, 'Scan interval time in msec between two samples (to be set once)'), (188, 'Scanned data with time label (can be readout event by event)'), (189, 'First-step offset current/voltage for valve when opening from 0%'), (190, 'Amount of runs of a piston prover (0 = stability check)'), (191, 'Minimum process time of a piston prover in 0.1 seconds'), (192, 'Leak rate piston prover'), (193, 'Sets instr. in info mode for 1 read-cycle to check available parameter options'), (194, 'Gives info about possible values of a mode in an array as result of ModeInfo req'), (195, 'Gives description about one of the mode options'), (196, 'Enables/disables options for calibration device (8 bits for 8 options)'), (197, 'Real mass flow in kg/min'), (198, 'Fieldbus address (top interface)'), (199, 'Configuration setting for interface to other bus-systems'), (200, 'Fieldbus baudrate (top interface)'), (201, 'Fieldbus diagnose string (top interface)'), (202, 'Number of vanes for use in a rotor meter'), (203, 'Fieldbus name (top interface)'), (204, 'measured value for direct reading (in capunits, max.= capacity)'), (205, 'setpoint: wanted value for direct reading (in capunits, max.= capacity)'), (206, 'Mass in g'), (207, 'Manufacturer Status register (64 diagnostic bits)'), (208, 'Manufacturer Warning register (64 diagnostic bits)'), (209, 'Manufacturer Error register (64 diagnostic bits)'), (210, 'Diagnostic history string (contains history of diag codes)'), (211, 'Diagnostic mode (0 = diagnostics off, 1 = diagnostics on)'), (212, 'Manufacturer Status enable (0-127 or 254 = disable all, 255 = enable all)'), (213, 'Analog measure output, zero adjust'), (214, 'Analog measure output, span adjust'), (215, 'Analog setpoint input, zero adjust'), (216, 'Analog setpoint input, span adjust'), (217, 'Sensor input, zero adjust'), (218, 'Sensor input, span adjust'), (219, 'Sensor temperature input, zero adjust'), (220, 'Sensor temperature input, span adjust'), (221, 'Sensor input filter adapt setting'), (222, 'Slope setpoint step. Setpoint step for the given slopetime'), (223, 'Number of samples for average filter'), (224, 'Actual accuracy in current unit'), (225, 'Lookup table for linearisation index (x and y direction)'), (226, 'Lookup table for linearisation x'), (227, 'Lookup table for linearisation y'), (228, 'Lookup table for linearisation at certain temperature index (z direction)'), (229, 'Lookup table for linearisation at certain temperature (z)'), (230, 'Maximum current/voltage for valve'), (231, 'Valve output mode selection'), (232, 'Valve open current/voltage correction (example: 0.96, Open = ValveOpen * 0.96)'), (233, 'Valve hold current/voltage at %0 setp (example: 0.8, Hold = ValveOpen * 0.8)'), (234, 'Valve slope time (seconds)'), (235, 'IFI data dump protocol communication string'), (236, 'Piston Prover information about used sensors'), (237, 'Fluidset properties'), (238, 'Lookup table unit type'), (239, 'Lookup table unit type name'), (240, 'Lookup table unit (unit LUTy)'), (241, 'Lookup table unit name'), (244, 'Capacity (readout) unit type temperature (°C)'), (245, 'Capacity (readout) unit type pressure (bar (a))'), (248, 'Formula type needed for conversion'), (249, 'Heat capacity (Cp) (sensor conditions)'), (250, 'Thermal conductivity (sensor conditions)'), (251, 'Dynamic viscosity (fluid conditions)'), (252, 'Standard mass flow in ln/min (20°C, 1.01325 bar (a) ) air or g/h H2O equivalent'), (253, 'Controller speed factor (gain)'), (254, 'Sensor code'), (255, 'Sensor configuration code'), (256, 'Restriction code'), (257, 'Restriction configuration code'), (258, 'Restriction NxP (proportional to air equivalent capacity of LFE)'), (259, 'Seals information (1st byte = other, 2nd = plunger seal)'), (260, 'Valve code'), (261, 'Valve configuration code'), (262, 'Instrument properties'), (263, 'Lookup table for frequency index'), (264, 'Lookup table for frequency frequency'), (265, 'Lookup table for frequency temperature'), (266, 'Lookup table for frequency density'), (267, 'Lookup table for frequency span adjust'), (268, 'Capacity (readout) unit index (extended unit table)'), (269, 'Actual density, measured by instrument'), (270, 'Measured restriction'), (271, 'Potmeter for sensor temperature compensation'), (272, 'Gain for sensor temperature compensation'), (273, 'Counter controller overrun correction'), (274, 'Counter controller gain'), (275, 'Sub fluid number'), (276, 'Temperature compensation factor'), (277, 'DSP register address'), (278, 'DSP register long'), (279, 'DSP register floating point'), (280, 'DSP register integer'), (281, 'Standard deviation'), (282, 'Measurement status'), (283, 'Measurement stop criteria'), (284, 'Measurement time out'), (285, 'Measurement maximum number of runs'), (286, 'Measurement minimum standard deviation'), (287, 'IO status for switches'), (288, 'Sensor bridge settings'), (289, 'Sensor bridge current (0..0.02A)'), (290, 'Sensor resistance (Ohm)'), (291, 'Sensor bridge voltage (0..11.5V)'), (292, 'Sensor name (based on sensor detection / sensor resistance)'), (293, 'Sensor temperature at calibration'), (294, 'Valve safe state'), (295, 'Counter unit type'), (296, 'Counter unit type name'), (297, 'Counter unit index (extended counter unit table)'), (298, 'Fieldbus select'), (299, 'Fieldbus medium id'), (300, 'Fieldbus2 mode'), (301, 'Fieldbus2 select'), (302, 'Fieldbus2 address (side interface)'), (303, 'Fieldbus2 baudrate (side interface)'), (304, 'Fieldbus2 medium id'), (305, 'Fieldbus2 diagnose string (side interface)'), (306, 'Fieldbus2 name (side interface)'), (307, 'PIO channel selection'), (308, 'PIO parameter connected to i/o channel (=process number*256+parameter number)'), (309, 'PIO input/output filter constant (0..1, 0=max, 1=min)'), (310, 'PIO parameter capacity, 0% value'), (311, 'PIO parameter capacity, 100% value'), (312, 'PIO i/o channel configuration selection'), (313, 'PIO analog input/output, zero adjust value'), (314, 'PIO analog input/output, span adjust factor'), (315, 'PIO max hardware capacity (max capacity value in i/o units, V, A, Hz, etc.)'), (316, 'PIO capacity set selection'), (317, 'PIO hardware capacity, 0% value (capacity value in i/o units, V, A, Hz, etc.)'), (318, 'PIO hardware capacity, 100% value (capacity value in i/o units, V, A, Hz, etc.)'), (319, 'Hardware platform identification number'), (320, 'Hardware platform sub identification number'), (321, 'Temporary volatile Fieldbus2 baudrate (side interface)'), (322, 'Monitor mode for setpoint'), (323, 'Special BHT parameter'), (324, 'Nominal sensor voltage (used for sensor temperature voltage compensation)'), (325, 'Sensor voltage compensation factor (used for sensor temperature voltage comp.)'), (326, 'PCB serial number (unique number or string written by PCB manufacturer)'), (327, 'Minimum measure time between sensors of a piston prover in 0.1 seconds'), (328, 'Fieldbus parity'), (329, 'Fieldbus2 parity (side interface)'), (330, 'Firmware identification number')])